## Slide 1
Pitfalls in Large-Scale Metabolomic Study of Human Plasma
Tomáš Čajka, Ph.D.
Metabolomics Seminar
December 4, 2013

## Slide 2
NIH: P20 PROJECT
Introduction
Metabolic pathways associated with atherosclerosis and cardiovascular risk factor
(RFA-HL-12-009 Anchoring Metabolomic Changes to Phenotype)
04/01/2012—03/31/2017

## Slide 3
ATHEROSCLEROSIS
Introduction
Myocardial infarction, stroke, and heart failure  the leading causes of death in the US
The disease primarily responsible for these major adverse cardiovascular events is atherosclerosis
  A gradual thickening of the artery wall due to the accumulation of cholesterol and triglycerides
Pepine CJ., Am J Cardiol 1998:82 Suppl 10A, 23S–27S

## Slide 4
Introduction
Atherosclerosis is an inflammatory response that results in hard plaques composed of three layers:
  The artheroma (soft, flaky portion composed of macrophage white blood cells and lipids)
  Underlying layers of cholesterol crystals
  An outer calcified layer in advanced lesions
ATHEROSCLEROSIS

## Slide 5
Introduction
A piece of “black” statistics
  For 66% of men and 47% of women, the first sign of atherosclerosis is myocardial infarction (heart attack) that occurs when the atherosclerotic lesion breaks free of the artery wall and blocks blood flow.
  Most cases of myocardial infarction occur when an atherosclerotic lesion causes less than 50% occlusion of the artery, yet we are currently unable to detect plaques smaller than 75% occlusion
Nissen SE, Nicholls SJ, Sipahi I, et al. (2006). “Effect of very high-intensity statin therapy on regression of coronary atherosclerosis: the ASTEROID trial”. JAMA295 (13): 1556–65
ATHEROSCLEROSIS

## Slide 6
NIH: P20 PROJECT
Introduction
Identification of novel pathways associated with atherosclerosis and its major adverse events (myocardial infarction, stroke, heart failure, or death), quantitative measures of atherosclerosis progression, and clinical response to statin therapy in humans
  Unbiased metabolomics approaches
  GeneBank samples (n = 1163)
  50 µL of plasma per sample
Study #1
Limitation for current GC and LC methods

## Slide 7
NIH: P20 PROJECT
Introduction
Intravascular Ultrasound (IVUS) Trial
  Every patient has quantitative measure of their coronary plaque volume, and has plaque progression/regression rate known since they had this procedure (IVUS) done on two occasions 18 months apart
  Unbiased metabolomics approaches
  Cleveland Clinic samples (n = 540)
Study #2

## Slide 8
SEPARATION–MS-BASED METABOLOMICS APPROACHES
Introduction
Triacylglycerols
Non-polar Vitamins
Sterols
Non-polar metabolites
Medium-polar metabolites
Polar metabolites
Very-polar metabolites
Fatty acids, PC, LPC
HILIC
RP-(U)HPLC
Amino acids
Organic acids
Carnitine
Monosaccharides
TMAO
(Reversed Phase Ultra-High Performance Liquid Chromatography)
(Hydrophilic Interaction Chromatography)
GC
(Gas Chromatography)

## Slide 9
SEPARATION–MS-BASED METABOLOMICS APPROACHES
Introduction
RP-(U)HPLC
(Reversed Phase Ultra-High Performance Liquid Chromatography)
Scope: Lipids
Mechanism of action in RP-LC of lipids is based on lipophilicity, which is governed by the carbon chain length and the number of double bonds
LPC, MAG
PC, SM, Cer, DAG
CE, TAG
Sandra, K. et al., J. Chromatogr. A, 2010, 1217, 4087

## Slide 10
SEPARATION–MS-BASED METABOLOMICS APPROACHES
Introduction
Scope: Primary metabolites
Derivatization required
Stability of the extracts max. 2 days
Some compounds undergo degradation
  Arginine  ornithine
  Creatinine  creatine
Formation of multiple peaks (due to derivatization)
GC
(Gas Chromatography)

## Slide 11
SEPARATION–MS-BASED METABOLOMICS APPROACHES
Introduction
Scope: Very-polar and polar metabolites
Complementary technique to RP-LC and GC (derivatization)
Less prone to thermal degradation of compounds as compared to GC
Longer equilibration time as compared to RP-LC
HILIC
(Hydrophilic Interaction Chromatography)

## Slide 12
GENERAL SCHEME OF ANALYSES

## Slide 13
Sample Extraction: SOP PROTOCOL
GC–TOFMS (primary metabolites)
30 µL plasma
+ 1000 µL ice-cold ACN:IPA:H2O
Vortexing (10 s)
Shaking (5 min; 4°C)
Re-suspending with 50% ACN 450 µL
Remove supernatant to a new Eppendorf tube
Drying: one tube only
(second tube as back-up)
Centrifugation (2 min; 14,000×g)
Transfer the upper phase to two separate tubes (450 µL/each tube)
Centrifugation (2 min; 14,000×g)
Drying
Derivatization
10 µL MeOX (40 mg/mL in pyridine)
Shaking (1.5 h; 30°C)
100 µL MSTFA + FAME mixture
Shaking (37°C)
Transfer content to glass vial with micro-insert
GC–EI-TOFMS analysis

## Slide 14
30 µL plasma
+ 1000 µL ice-cold ACN:IPA:H2O
Vortexing (10 s)
Shaking (5 min; 4°C)
Re-suspending with 50% ACN 500 µL
Remove supernatant to a new Eppendorf tube
Drying: one tube only
(second tube as back-up; third tube for HILIC)
Centrifugation (2 min; 14,000×g)
Transfer the upper phase to three separate tubes (300 µL/each tube)
Centrifugation (2 min; 14,000×g)
Drying
Derivatization
10 µL MeOX (40 mg/mL in pyridine)
Shaking (1.5 h; 30°C)
60 µL MSTFA + FAME mixture
Shaking (37°C)
Transfer content to glass vial with micro-insert
GC–EI-TOFMS analysis
Sample Extraction: P20 PROTOCOL
GC–TOFMS (primary metabolites)

## Slide 15
BinBase
Example of GC–EI-TOFMS profile of human plasma (primary metabolites)
Completely automated analysis tool (database system) for the annotations of GC–EI-TOFMS data
DATA PROCESSING

## Slide 16
Sample Extraction: SOP PROTOCOL
UHPLC–QTOFMS (LIPIDS)
20 µL plasma
+ 225 µL ice-cold MeOH (containing QC mix)
Vortexing 10 s
Add 750 µL ice-cold MTBE (containing CE(22:1))
Vortexing (10 s)
Shaking (6 min; 4°C)
Add distilled water 188 µL
Vortexing (20 s)
Centrifugation (2 min; 14,000×g)
Transfer the upper phase to two separate tubes (350 µL/each tube)
Drying: one tube only
(second tube as back-up)
Reconstitution: 108.6 µL MeOH:toluene (9:1, v/v) (containing CUDA IS)
Centrifugation (2 min; 14,000×g)
Transfer extracts to two separate glass amber vials with micro-insert
UHPLC–QTOFMS analysis
1.67 µL injection ESI(+) Agilent 6530
5 µL injection ESI(–) Agilent 6550

## Slide 17
TRANSFER OF THE METHOD FROM THE AGILENT 6530 TO 6550
8
x10
0
0.05
0.1
0.15
0.2
0.25
0.3
0.35
0.4
0.45
0.5
0.55
0.6
0.65
0.7
0.75
0.8
0.85
0.9
0.95
1
1.05
Counts vs. Acquisition Time (min)
1
2
3
4
5
6
7
8
9
10
11
12
Sample: NIST plasma
Agilent 6550 ESI(+)
Plasma equivalent injected: ~331 nL (TEDDY PROTOCOL v. 1)
Many saturated ions

## Slide 18
10 µL plasma

225 µL MeOH
+ 750 µL MTBE
+ 188 µL H2O

350 µL aliquot

130 µL resusp.
in MeOH:toluene (9:1, v/v)

3 µL injection
LC–(ESI+/–)-QTOFMS 6550
Plasma equivalent injected: ~83 nL
(25%)
Method A
Method B
Method C
10 µL plasma

225 µL MeOH
+ 750 µL MTBE
+ 188 µL H2O

200 µL aliquot

150 µL resusp.
in MeOH:toluene (9:1, v/v)

3 µL injection
LC–(ESI+/–)-QTOFMS 6550
Plasma equivalent injected: ~41 nL
(12%)
10 µL plasma

225 µL MeOH
+ 750 µL MTBE
+ 188 µL H2O

100 µL aliquot

200 µL resusp.
in MeOH:toluene (9:1, v/v)

3 µL injection
LC–(ESI+/–)-QTOFMS 6550
Plasma equivalent injected: ~15 nL
(5%)
Method SOP
20 µL plasma

225 µL MeOH
+ 750 µL MTBE
+ 188 µL H2O

350 µL aliquot

65 µL resusp.
in MeOH:toluene (9:1, v/v)

3 µL injection
LC–ESI(+)-QTOFMS 6530
Plasma equivalent injected: ~331 nL
(100%)
TRANSFER OF THE METHOD FROM THE AGILENT 6530 TO 6550
Agilent 6530
Agilent 6550

## Slide 19
PC(34:2)
m/z 758.569
Method A
Method B
Method C
TAG(52:3)
m/z 874.7887
LPC(16:0)
m/z 496.3398
1 µL
2 µL
3 µL
4 µL
5 µL
1 µL; dil. 2x
1 µL
2 µL
3 µL
4 µL
5 µL
1 µL; dil. 2x
1 µL
2 µL
3 µL
4 µL
5 µL
1 µL; dil. 2x
1 µL; dil. 4x
1 µL; dil. 8x
1 µL; dil. 10x
Method SOP
1 µL
2 µL
3 µL
4 µL
5 µL
TRANSFER OF THE METHOD FROM THE AGILENT 6530 TO 6550
LINEARITY

Compromise:
Linearity vs. number of features

## Slide 20
TRANSFER OF THE METHOD FROM THE AGILENT 6530 TO 6550
8
x10
0
0.05
0.1
0.15
0.2
0.25
0.3
0.35
0.4
0.45
0.5
0.55
0.6
0.65
0.7
0.75
0.8
0.85
0.9
0.95
1
1.05
Counts vs. Acquisition Time (min)
1
2
3
4
5
6
7
8
9
10
11
12
Sample: NIST plasma
Agilent 6550 ESI(+)
Plasma equivalent injected: ~41 nL (P20 PROTOCOL)
Plasma equivalent injected: ~331 nL (SOP PROTOCOL v. 1)
Plasma equivalent injected: ~110 nL (SOP PROTOCOL v. 2)

## Slide 21
10 µL plasma
+ 225 µL ice-cold MeOH (containing QC mix)
Vortexing 10 s
Add 750 µL ice-cold MTBE (containing CE(22:1))
Vortexing (10 s)
Shaking (6 min; 4°C)
Add distilled water 188 µL
Vortexing (20 s)
Centrifugation (2 min; 14,000×g)
Transfer the upper phase to two separate tubes (200 µL/each tube)
Drying: one tube only
(second tube as back-up)
Reconstitution: 150 µL MeOH:toluene (9:1, v/v) (containing CUDA IS)
Centrifugation (2 min; 14,000×g)
Transfer extracts to two separate glass amber vials with micro-insert
UHPLC–QTOFMS analysis
3 µL injection ESI(+) Agilent 6550
3 µL injection ESI(–) Agilent 6550
Sample Extraction: P20 PROTOCOL
UHPLC–QTOFMS (LIPIDS)

## Slide 22
Separation of Lipids:
Ultra-High Performance LC
CSH (Charged Surface Hybrid) Technology
Column: Acquity UPLC CSH C18 (100  2.1 mm; 1.7 µm)
  Good retention and separation of lipid classes (PC, lysoPC, PE, PS, TG, ceramides)
  Peak width: 8–17 s
  Excellent repeatability of retention times
Mobile phase A: 60:40 ACN:H2O + 10 mM AmF + 0.1% FA
Mobile phase B: 90:10 IPA:ACN + 10 mM AmF + 0.1% FA
Temperature: 65°C
Flow-rate: 0.6 mL/min
Sample temperature: 4°C

## Slide 23
CSH (Charged Surface Hybrid) Technology
Based on Waters Ethylene Bridged Hybrid (BEH) particle technology
CSH particles incorporate a low level surface charge
Designed to improve sample loadability and peak asymmetry in low-ionic-strength mobile phases, while maintaining the mechanical and chemical stability inherent in BEH particle technology
http://www.waters.com
The advantages of CSH Technology include:
Superior peak shape for basic compounds
Increased loading capacity
Rapid column equilibration after changing mobile-phase pH
Improved batch-to-batch reproducibility
Exceptional stability at low and high pH

## Slide 24
1
2
3
4
5
6
7
8
9
10
11
12
13
1
2
3
4
5
6
7
8
9
10
11
12
13
14
ESI(+)
ESI(–)
Gradient:
0 min 15% (B)
0–2 min 30% (B)
2–2.5 min 48% (B)
2.5–11 min 82% (B)
11–11.5 min 99% (B)
11.5–12 min 99% (B)
12–12.1 min 15% (B)
12.1–15 min 15% (B)
Gradient:
0 min 15% (B)
0–2 min 30% (B)
2–2.5 min 48% (B)
2.5–9.5 min 76% (B)
9.5–9.6 min 99% (B)
9.6–10.5 min 99% (B)
10.5–10.6 min 15% (B)
10.6–13.5 min 15% (B)
15 min
run-to-run
cycle
13.5 min
run-to-run
cycle
CE, TG
PC, SM, Cer, DG
LPC, LPE
FA
SM, PE, PG, Cer
LPC, LPE
Separation of Lipids:
Ultra-High Performance LC

## Slide 25
QTOFMS DETECTION
Agilent 6530
ESI(+)
Agilent 6550
ESI(–)
Agilent 6550
ESI(+)
Agilent 6550
ESI(–)
20 µL
plasma
10 µL
plasma
Extraction procedure
Extraction procedure
SOP PROTOCOL
P20 PROTOCOL
Mass resolving power: 10,000 FWHM
Mass resolving power: 20,000 FWHM
4 working days / 1 batch 400 injections
8 working days / 1 batch 400 injections

## Slide 26
MassHunter — Quantitative Analysis
Example of LC–ESI(+/–)-QTOFMS profile of human plasma (lipids)
Target approach for reporting intensities (or concentrations of analytes)
IS used for RT correction (as well as quantification)
DATA PROCESSING
Identified compounds
Unknowns
compound name
quantification ion
peak height
retention time
RT_MZ ID
elemental composition
InChIKey
conc. of IS

## Slide 27
OVERVIEW OF IDENTIFIED LIPIDS
ESI(+)
ESI(–)

## Slide 28
1
2
3
4
5
6
7
8
9
10
11
12
13
14
1
2
3
4
5
6
7
8
9
10
11
12
13
14
1
2
3
4
5
6
7
8
9
10
11
12
13
1
2
3
4
5
6
7
8
9
10
11
12
(1) CUDA *)
(2) Sphingosine (d17:1)
(3) LysoPE(17:1)
(4) LysoPC(17:0)
(5) MG(17:0/0:0/0:0)
(11) C17 Ceramide
(12) PE(17:0/17:0)
(13) d5-TG(17:0/17:1/17:0)
(14) CE(22:1)
(6) DG(18:1/2:0/0:0)
(7) PC(12:0/13:0)
(8) DG(12:0/12:0/0:0)
(9) d7-Cholesterol
(10) SM(17:0)
(1) CUDA *)
(2) LysoPE(17:1)
(3) LysoPC(17:0)
(4) d3-Palmitic acid
(5) SM(17:0)
(6) PG(17:0/17:0)
(7) PE(17:0/17:0)
*) IS used for quality control (injection)
ESI(+)
TIC
EIC
(internal standards)
1
2
3
4
5
6
7
8
9
10
11
13
14
12
ESI(–)
TIC
EIC
(internal standards)
1
2
3
4
5
6
7
A mixture of odd chain and deuterated lipid internal standards
QUANTIFICATION OF LIPIDS

## Slide 29
30 µL plasma
+ 1000 µL ice-cold ACN:IPA:H2O
Vortexing (10 s)
Shaking (5 min; 4°C)
Re-suspending with
80% ACN 60 µL
(containing a mixture of internal standards)
Drying: one tube only
(second tube for GC–TOFMS; third tube as a backup)
Centrifugation (2 min; 14,000×g)
Transfer the upper phase to three separate tubes (300 µL/each tube)
Centrifugation (10 min; 16,000×g)
Transfer 50 µL to glass vial with micro-insert
HILIC–ESI(+)-QTOFMS analysis
Sample Extraction: P20 PROTOCOL
HILIC–QTOFMS (POLAR METABOLITES)

## Slide 30
Separation of Polar Compounds:
Hydrophilic Interaction Chromatography (HILIC)
Column: ACQUITY UPLC BEH Amide Column (100  2.1 mm; 1.7 µm)
  Chemically stable, trifunctionally-bonded amide phase
  Good retention and separation of polar metabolites (TMAO, betaine, amino acids)
  Peak width: 6–15 s
Mobile phase A: 10 mM AmF + 0.125% FA
Mobile phase B: 10 mM AmF + 0.125% FA 95:5 ACN:H2O
Temperature: 45°C
Flow-rate: 0.4 mL/min
Sample temperature: 4°C

## Slide 31
Separation of Polar Compounds:
HILIC
(1) Choline-D9
(2) TMAO-D9
(3) Betaine-D9
(4) L-Carnitine-D3
(5) Creatine-D3
(6) L-Arginine-15N2
(1)
(2)
(6)
(5)
(4)
(3)
TIC
TripleTOF (AB SCIEX)
Acquisition
(1) Full Mass Spectra (m/z 55–1000)
(2) MRM for selected metabolites (TMAO, betaine)
Gradient
0 min 100% B
0 min  2 min 100% B
2 min  7.7 min 70% B
7.7 min  9.5 min 40% B
9.5 min  10.25 min 30% B
10.25 min  12.75 100% B
16.75 min 100% B
(7) Glutamine-D5
(8) -Butyrobetaine (labeled)
(9) trans-Crotonobetaine (labeled)

## Slide 32
DISCOVERING PITFALLS
Preparation of QC mix
Saturation of the MCP detector and electrospray
High-fat plasma samples
Normalization strategies
Comparison of ESI(+) and ESI(–) quantitative data for lipids detected in both modes

## Slide 33
PREPARATION OF QC MIX
Possible pitfalls
Stability of standards
  Store all stock solutions at –20°C
Precipitation during storage at –20°C
  Allow the stock standards to equilibrate at room temperature (min. 30 min)
  Solubility of some standards needs to be enhanced by sonication
Evaporation of solvents from stock solutions
  Use of Certan capillary bottle (Kapillarflasche): 4.5 mL or 1.5 mL
    Vial with a narrow 0.8 mm hole
    Secure storage, minimal evaporation

## Slide 34
PREPARATION OF QC MIX
Possible pitfalls
Preparation of stock solutions and “QC mix” extraction solvents
  Compare the new “QC mix” solution with the old one by LC–MS analysis: the deviation should be <10%
  Example: IS-PC(12:0/13:0)
Batch #1 vs. Batch #2
Batch #2 vs. Batch #3
Batch #3 vs. Batch #4
Peak height
10%

## Slide 35
SATURATION OF THE MCP DETECTOR AND ELECTROSPRAY
Saturation of the MCP detector and/or ADC
  Multiple ions with the same m/z arrive at the detector at the same time.
  Only a part of ions is detected/recorded.
  Isotope profile is distorted.
Order of magnitude
0
2
4
6
8
Ion intensity
104
105
106
107
108
Saturation of MCP/ADC
Linear range

## Slide 36
7
x10
0
0.05
0.1
0.15
0.2
0.25
0.3
0.35
0.4
0.45
0.5
0.55
0.6
0.65
0.7
0.75
0.8
0.85
0.9
0.95
1
1.05
1.1
* 848.7787
* 849.7802
850.7805
851.7822
Counts vs. Mass-to-Charge (m/z)
849
849.5
850
850.5
851
851.5
852
849.0
849.5
850.0
850.5
851.0
851.5
852.0
m/z
0
10
20
30
40
50
60
70
80
90
100
Relative Abundance
848.7702
849.7735
850.7768
851.7798
SATURATION OF THE MCP DETECTOR
Example: TG(50:2), m/z 848.7
Isotope profile: saturation
Theoretical isotope profile
Saturated ions

## Slide 37
SATURATION OF THE MCP DETECTOR
7
x10
0
0.05
0.1
0.15
0.2
0.25
0.3
0.35
0.4
0.45
0.5
0.55
0.6
0.65
0.7
0.75
0.8
0.85
0.9
0.95
1
1.05
1.1
Counts vs. Mass-to-Charge (m/z)
849
849.5
850
850.5
851


–30%
–43%
[M+NH4]+
[M+1+NH4]+
[M+2+NH4]+
[M+3+NH4]+
Theoretical isotope profile
[M+NH4]+ 100%
[M+1+NH4]+ 59.1%
[M+2+NH4]+ 18.4%
Correction factor:
100/18.4 = 5.43
Isotope profile: saturation
Theoretical isotope profile

## Slide 38
SATURATION OF THE MCP DETECTOR
Saturation observed for 7 PCs and 7 TGs
The “true” intensity is obtained by multiplying the intensity of [M+2+H]+ or [M+2+NH4]+ ion by an appropriate correction factor

## Slide 39
SATURATION OF THE MCP DETECTOR
ESI(+)
(Uncorrected) Concentration (µM)
Calculated using [M+H]+ intensity
ESI(+)
(Corrected) Concentration (µM)
Calculated using [M+2+H]+ intensity
and correction factor
Range: 74–312 µM
Range: 99–504 µM
Comments:
In batch #1 (n = 300) the concentrations obtained in ESI(+) were around 43% lower as compared to values obtained after correction
Example: PC(34:2)

## Slide 40
SATURATION OF THE MCP DETECTOR AND ELECTROSPRAY
Saturation of electrospray
  Related to formation of droplets during the ionization process
   flow rate  linearity
Order of magnitude
0
2
4
6
8
Ion intensity
104
105
106
107
108
Saturation of ESI
Linear range
Order of magnitude
0
2
4
6
8
Ion intensity
104
105
106
107
108
Saturation of ESI
Linear range
Analyte 1
Analyte 2
No saturation, no suppression
Analyte 1: ESI saturation
Analyte 2: Suppression

## Slide 41
SATURATION OF ELECTROSPRAY
7
x10
0
1
2
3
4
5
6
7
5
x10
0.2
0.4
0.6
0.8
1
1.2
1.4
Counts vs. Acquisition Time (min)
1
2
3
4
5
6
7
8
9
10
11
12
13
Reference ion
m/z 922.009
TIC
Ideal signal intensity (no ESI saturation)
ESI saturation regions
Example: High-fat plasma sample
Real signal intensity
Notes: B1_SA0216

## Slide 42
SATURATION OF ELECTROSPRAY
Solution: re-analysis of samples after dilution
Question #1: What should be a criterion for selection of sample for dilution and re-analysis? Difficult to answer…
Question #2: Do we need “more” accurate results for those few most abundant lipids? As PCA shows, the problem is related mainly to “outliers”, that tend to be even more remote
A (S)
A (NS)
B (S)
B (NS)
C (S)
C (NS)
Injection: 3 µL (with suppression)
Injection: 1 µL (without suppression)
Notes: B1_SA0216

## Slide 43
NORMALIZATION STRATEGIES
How can we ensure metabolite data integrity and quantitative robustness in large-scale metabolomics studies?
blank
QC
NIST
blank
QC
normalize to
normalize to
samples
samples
Randomization, block design
Blank and machine quality controls
Reference (QC) and reference standard sample (NIST blood)
Use of multiple internal standards (if applicable)
Multiple internal standards

## Slide 44
NORMALIZATION STRATEGIES
Constant row sum using mTIC or TIC
  CRS = each variable is divided by the sum of all variables for each sample
  mTIC = intensity of all identified metabolites (=knowns)
  TIC = intensity of all metabolites (=knowns + unknowns)
Single point quantification using internal standard
  Expressed in concentration units (µM)
  For uknowns the concentrations are calculated using the nearest internal standards
Batch normalization
  Using QC samples in all batches and correction factors

## Slide 45
NORMALIZATION STRATEGIES
Data set: ESI(+)
Samples + QC + NIST
Knowns (n = 154)
Data transformation:
  CRS (mTIC)
  log
  Pareto
PCA
Score plot

## Slide 46
CE
TG (with unsaturated FA)
TG (with saturated FA)
Ceramides
SM
LPC
DG
PC (with saturated FA)
PC (with unsaturated FA)
Data set: ESI(+)
Samples only
Knowns (n = 154)
Data transformation:
  CRS (mTIC)
  log
  Pareto
NORMALIZATION STRATEGIES
PCA
Score plot
PCA
Loading plot

## Slide 47
NORMALIZATION STRATEGIES
Constant row sum using mTIC or TIC
  CRS = each variable is divided by the sum of all variables for each sample
  mTIC = intensity of all identified metabolites (=knowns)
  TIC = intensity of all metabolites (=knowns + unknowns)
Single point quantification using internal standard
  Expressed in concentration units (µM)
  For uknowns the concentrations are calculated using the nearest internal standards
Batch normalization
  Using QC samples in all batches and correction factors

## Slide 48
NORMALIZATION STRATEGIES
Single point quantification using internal standard
Batch normalization
  Calculation of median Cm or median ratio (hm/hIS) for each metabolite in each batch of QC
  Calculation of correction factor (from QC samples)
    Fm;2,1 = med(batch 2) / med(batch 1)
    Fm;3,1 = med(batch 3) / med(batch 1)
    Fm;4,1 = med(batch 4) / med(batch 1)
  Application of correction factors to batches 2–4 (to all samples and QC/NIST samples)
    Cm, batch 2, corr = Cm, batch 2, uncorr / Fm;2,1
    Cm, batch 3, corr = Cm, batch 3, uncorr / Fm;3,1
    Cm, batch 4, corr = Cm, batch 4, uncorr / Fm;4,1

## Slide 49
NORMALIZATION STRATEGIES
Data set: ESI(+)
Samples + QC + NIST
Knowns (n = 154)
Data transformation:
  IS correction
  log
  Pareto
PCA
Score plot

## Slide 50
Data set: ESI(+)
Samples + QC + NIST
Knowns (n = 154)
Data transformation:
  IS correction
  batch normalization
  log
  Pareto
NORMALIZATION STRATEGIES
PCA
Score plot

## Slide 51
CE
TG (with unsaturated FA)
TG (with saturated FA)
Ceramides
PC
SM
LPC
DG
Data set: ESI(+)
Samples only
Knowns (n = 154)
Data transformation:
  IS correction
  batch normalization
  log
  Pareto
NORMALIZATION STRATEGIES
PCA
Score plot
PCA
Loading plot

## Slide 52
Data set: ESI(–)
Samples + QC + NIST
Knowns (n = 152)
Data transformation:
  IS correction
  log
  Pareto
NORMALIZATION STRATEGIES
PCA
Score plot

## Slide 53
Data set: ESI(–)
Samples + QC + NIST
Knowns (n = 152)
Data transformation:
  IS correction
  batch normalization
  log
  Pareto
NORMALIZATION STRATEGIES
PCA
Score plot

## Slide 54
Data set: ESI(–)
Samples only
Knowns (n = 152)
Data transformation:
  IS correction
  batch normalization
  log
  Pareto
NORMALIZATION STRATEGIES
PC
Plasmenyl-PC
Plasmenyl-PE
Ceramides
Free FA
PS
PE
SM
LPC
PCA
Loading plot
PCA
Score plot

## Slide 55
COMPARISON OF QUANTITATIVE RESULTS: ESI(+) and ESI(–)
Example #1: High abundant PC(34:2)
ESI(+)
(Uncorrected) Concentration (µM)
Calculated using [M+H]+ intensity
ESI(+)
(Corrected) Concentration (µM)
Calculated using [M+2+H]+ intensity
and correction factor
ESI(–)
Concentration (µM)
Comments:
In batch #1 (n = 300) the corrected concentrations obtained in ESI(+) were around 34% lower as compared to ESI(–)
For comparison, 62% lower if concentrations were not corrected in ESI(+)
Possible explanation is ion suppression during ESI(+)
Range: 149–764 µM
Range: 74–312 µM
Range: 99–504 µM

## Slide 56
COMPARISON OF QUANTITATIVE RESULTS: ESI(+) and ESI(–)
Example #2: Low abundant PC(36:1)
ESI(+)
Concentration (µM)
ESI(–)
Concentration (µM)
Range: 3.9–80 µM
Range: 4.8–126 µM
Comments:
In batch #1 (n = 300) the concentrations obtained in ESI(+) were around 18% lower as compared to ESI(–)
Possible explanation is ion suppression during ESI(+)

## Slide 57
QUALITY CONTROL:
TUNING OF THE INSTRUMENT
3.0M
9.0M
Standard Tune #1
Initial Tune +
Standard Tune #2
(1) Check always the ion intensity of calibration solution
(2) Lower ion intensity of calibration solution indicates decrease of signal intensity
(3) Recommendations: Clean the ion source, conduct initial tune (>1 h), follow standard tune (20 min)
7
x10
0
0.5
1
1.5
2
2.5
3
3.5
4
4.5
5
Counts vs. Acquisition Time (min)
1
2
3
4
5
6
7
8
9
10
11
12
 Sample run after Standard tune #1
 Sample run after Initial tune + Standard tune #2

## Slide 58
QUALITY CONTROL:
Common Contaminants in LC–MS
Polysiloxanes
  Source: septa
PEG
  Source: solvents, tubing
+74 series
+22 and +44 series

## Slide 59
QUALITY CONTROL:
Common Contaminants in LC–MS
Phthalates
  Source: plastics
Alkanes
  Source: solvent
“No injection” TIC profile with “pure” solvents used
“No injection” TIC profile with contaminated
solvent (IPA) used
Di-n-octyl phthalate
m/z 391.3 [M+H]+
m/z 413.3 [M+Na]+
(1) Check always the background noise for 100% mobile phase A and 100% mobile phase B prior the sequence.
(2) Check always the quality of the solvents when changing manufacturer and/or lot.
(3) Note that “LC–MS grade” does not mean “totally free” of ALL contaminants.

## Slide 60
PUBLICATIONS
Analysis of Complex Lipid Mixtures in Biological Systems Using Liquid Chromatography–Mass Spectrometry
• Review paper (TrAC-Trends in Analytical Chemistry)
Advantages and Limitations of (Ultra-)High Resolution Mass Spectrometry in the Analysis of Complex Lipids in Human Plasma
• Comparison of Citius, Agilent 6550
Discovering Pitfalls in Large-Scale Metabolomic Study of Human Plasma
• Based on the analysis of ca. 1160 samples (P20)
Pathway Phenotyping Study
• Single-blinded, placebo-controlled, postprandial, crossover in triplicate study designed to determine the effects of a high gamma-linolenic acid fat (GF) challenge compared with a high-monounsaturated fatty acid control fat (CF) challenge on plasma fatty acid composition
• n = 210 samples (7 participants; 2 dietary challenges (GF, CF); 3 different days (30 day wash-out period); blood samples at 0–2–4–6–8 hr); 7 x 2 x 3 x 5 = 210
• Lipids Analysis using UHPLC–QTOFMS in ESI(+) and ESI(–) (Agilent 6550)
Standardization in LC–MS-based Metabolomics
• Comparison of instrumental platforms (accent to quantitative aspect/normalization in metabolomics)
• n = 126 samples (out of 210 samples from the Pathway phenotyping study) (7 participants; 2 dietary challenges (GF, CF); 3 different days (30 day wash-out period); blood samples at 0–2–4 hr); 7 x 2 x 3 x 3 = 126
• Analysis in ESI(+) only
• Lipids (CSH) and polar compounds (HILIC)
• Instrumental platforms: Agilent 6550, TripleTOF, Citius, Q-Exactive

## Slide 61
Acknowledgements
Brian DeFelice
Carlos Leon
Alex Chassy
William Wikoff
Thank you for your attention…
tcajka@ucdavis.edu
St. Barbara's Church, Kutná Hora